## ENTRY
Pre-Measurement Calibration Check
Water Quality Instrument Record (SOP v1.0.0)
CA Department of Water Resources
Division of Environmental Services, Suisun Marsh Branch
[8 empty] | Sonde ID |  |  | SM-26
Discrete | x | Real Time |  |  | Performed by (Last Name) | [6 empty] | Pluton |  |  | Calibration Date |  | 2020-06-23
Site/Run to be used on |  |  |  | LAB | [9 empty] | Calibration Time (PST) |  |  | 730
Attachments?
Verification Instrument |  |  | Serial Number |  |  |  | Calibration Due Date |  |  |  |  | Calibration Standard |  |  | Lot Number |  | Exp. Date
[12 empty] | Turbidity Standard |  |  | 21833 |  | 2021-01-15
Thermometer |  |  | 191957329 |  |  |  | 2021-02-13 |  |  |  |  | Conductance Standard |  |  | 20700 |  | 2021-11-06
Turbidimeter |  |  | 2017080C0003 |  |  |  | 2020-07-05 |  |  |  |  | Conductance Check |  |  | 21860 |  | 2021-01-28
Barometer |  |  | 261829 | [8 empty] | 7 pH Standard |  |  | A0037 |  | 2022-02-01
addl. instrument | [11 empty] | 4 pH Standard
addl. instrument | [11 empty] | 10 pH Standard |  |  | A0016 |  | 2021-01-01
Pre-Measurement Calibration
To be completed no more than 72 hours prior to measurement. Follow current standard operating procedures.
Parameter | [5 empty] | Standard |  |  | Pre-Cal |  |  | Post-Cal | Passing Criteria |  |  | Additional Info
Specific Conductance (μS/cm) Dry | [5 empty] | 0 |  |  | -0.1 |  |  |  | < 2 μS/cm
Temperature (°C) in Water | [5 empty] | 22.08 |  |  | 21.94 |  |  |  | ≤ ± 0.20 °C
Chlorophyll (RFU) in DIW | [5 empty] | 0 |  |  | 0.01 |  |  | 0 | < 0.10 RFU
Chlorophyll (μg/L) in DIW | [5 empty] | 0 |  |  | 0.03 |  |  | 0 | < 0.10 μg/L
Blue Green Algae (RFU) in DIW | [12 empty] | < 0.10 RFU
Blue Green Algae (μg/L) in DIW | [12 empty] | < 0.10 μg/L
Turbidity (NTU) in DIW | [5 empty] | 0 |  |  | -0.8 |  |  | 0 | ≤ ± 0.5 NTU
Turbidity (NTU) in Std | [5 empty] | 100 |  |  | 103.97 |  |  | 100 | ≤ ± 5%
Specific Conductance (μS/cm) in Std | [5 empty] | 17500 |  |  | 17630 |  |  | 17500 | ≤ ± 3% |  |  | Cell constant: |  | 5.2
pH 7 (units) | [5 empty] | 7.01 |  |  | 7 |  |  | 7.01 | ≤ ± 0.2 units |  |  | -18.9 |  |  |  | mV
pH 4 (units) | [12 empty] | ≤ ± 0.2 units | [6 empty] | mV
pH 10 (units) | [5 empty] | 10.04 |  |  | 10.04 |  |  | 10.04 | ≤ ± 0.2 units |  |  | -192.6 |  |  |  | mV
Dissolved Oxygen (% Sat) | [5 empty] | 99.6184 |  |  | 100.4 |  |  | 99.4 | ≤ ± 5% |  |  | Baro. Pres. (mmHg): |  |  | 757.1
Dissolved Oxygen (mg/L) | [5 empty] | 9.0068 |  |  | 9.06 |  |  | 9 | ≤ ± 0.3 mg/L |  |  | Temp. (°C) in bucket : |  |  | 20.28
Dissolved Oxygen (Gain) | [11 empty] | 1.05 |  |  |  | Ideal range 0.87 - 1.25
Specific Conductance (μS/cm) Check | [5 empty] | 6668 | [5 empty] | 6611 | ≤ ± 3%
Delta slope (pH 7 mV - pH 10 mV) | [5 empty] | 173.7 |  |  |  | Ideal range 160 - 180 mV, replace module if slope ≤ 155 mV
Comments
Calibrate depth to 0 ft | [6 empty] | x | Day of deployment | [7 empty] | Verification Sonde
Verify Sufficient battery voltage | [6 empty] | x | Verify DO %sat ≤ 5 % | [6 empty] | x
Verify wiper is parked correctly and working | [6 empty] | x | Verify parameters & reporting | [6 empty] | x
Verify or set date & time (time.gov) | [6 empty] | x | Sonde started & ready for measurement | [6 empty] | x
Wiper activated during calibration | [6 empty] | x

## RESULT
cdec_code | interval_name | analyte_name | unit_name | rank_name | reading_type_name | time | value
LAB | Pre-measurement calibration | Specific Conductance | μS/cm | Standard: Specific Conductance - Dry | a. Calibration Standard | 2020-06-23 07:30:00 | 0
LAB | Pre-measurement calibration | Specific Conductance | μS/cm | Standard: Specific Conductance - Dry | b. Pre-calibration Reading | 2020-06-23 07:30:00 | 0
LAB | Pre-measurement calibration | Specific Conductance | μS/cm | Standard: Specific Conductance - Standard Solution | a. Calibration Standard | 2020-06-23 07:30:00 | 17500
LAB | Pre-measurement calibration | Specific Conductance | μS/cm | Standard: Specific Conductance - Standard Solution | b. Pre-calibration Reading | 2020-06-23 07:30:00 | 17630
LAB | Pre-measurement calibration | Specific Conductance | μS/cm | Standard: Specific Conductance - Standard Solution | c. Post-calibration Reading | 2020-06-23 07:30:00 | 17500
LAB | Pre-measurement calibration | Specific Conductance | μS/cm | Standard: Specific Conductance - Standard Check | a. Calibration Standard | 2020-06-23 07:30:00 | 6668
LAB | Pre-measurement calibration | Specific Conductance | μS/cm | Standard: Specific Conductance - Standard Check | c. Post-calibration Reading | 2020-06-23 07:30:00 | 6611
LAB | Pre-measurement calibration | Temperature | °C | Standard: Temperature - Water Bath | a. Calibration Standard | 2020-06-23 07:30:00 | 22.08
LAB | Pre-measurement calibration | Temperature | °C | Standard: Temperature - Water Bath | b. Pre-calibration Reading | 2020-06-23 07:30:00 | 21.94
LAB | Pre-measurement calibration | Chlorophyll | RFU | Standard: DIW | a. Calibration Standard | 2020-06-23 07:30:00 | 0.00
LAB | Pre-measurement calibration | Chlorophyll | RFU | Standard: DIW | b. Pre-calibration Reading | 2020-06-23 07:30:00 | 0.01
LAB | Pre-measurement calibration | Chlorophyll | RFU | Standard: DIW | c. Post-calibration Reading | 2020-06-23 07:30:00 | 0.00
LAB | Pre-measurement calibration | Chlorophyll | μg/L | Standard: DIW | a. Calibration Standard | 2020-06-23 07:30:00 | 0.00
LAB | Pre-measurement calibration | Chlorophyll | μg/L | Standard: DIW | b. Pre-calibration Reading | 2020-06-23 07:30:00 | 0.03
LAB | Pre-measurement calibration | Chlorophyll | μg/L | Standard: DIW | c. Post-calibration Reading | 2020-06-23 07:30:00 | 0.00
LAB | Pre-measurement calibration | Blue green algae | RFU | Standard: DIW | a. Calibration Standard | 2020-06-23 07:30:00
LAB | Pre-measurement calibration | Blue green algae | RFU | Standard: DIW | b. Pre-calibration Reading | 2020-06-23 07:30:00
LAB | Pre-measurement calibration | Blue green algae | RFU | Standard: DIW | c. Post-calibration Reading | 2020-06-23 07:30:00
LAB | Pre-measurement calibration | Blue green algae | μg/L | Standard: DIW | a. Calibration Standard | 2020-06-23 07:30:00
LAB | Pre-measurement calibration | Blue green algae | μg/L | Standard: DIW | b. Pre-calibration Reading | 2020-06-23 07:30:00
LAB | Pre-measurement calibration | Blue green algae | μg/L | Standard: DIW | c. Post-calibration Reading | 2020-06-23 07:30:00
LAB | Pre-measurement calibration | Turbidity | NTU | Standard: DIW | a. Calibration Standard | 2020-06-23 07:30:00 | 0.00
LAB | Pre-measurement calibration | Turbidity | NTU | Standard: DIW | b. Pre-calibration Reading | 2020-06-23 07:30:00 | -0.80
LAB | Pre-measurement calibration | Turbidity | NTU | Standard: DIW | c. Post-calibration Reading | 2020-06-23 07:30:00 | 0.00
LAB | Pre-measurement calibration | Turbidity | NTU | Standard: Turbidity - High | a. Calibration Standard | 2020-06-23 07:30:00 | 100.00
LAB | Pre-measurement calibration | Turbidity | NTU | Standard: Turbidity - High | b. Pre-calibration Reading | 2020-06-23 07:30:00 | 103.97
LAB | Pre-measurement calibration | Turbidity | NTU | Standard: Turbidity - High | c. Post-calibration Reading | 2020-06-23 07:30:00 | 100.00
LAB | Pre-measurement calibration | pH | pH Units | Standard: pH - 4 | a. Calibration Standard | 2020-06-23 07:30:00
LAB | Pre-measurement calibration | pH | pH Units | Standard: pH - 4 | b. Pre-calibration Reading | 2020-06-23 07:30:00
LAB | Pre-measurement calibration | pH | pH Units | Standard: pH - 4 | c. Post-calibration Reading | 2020-06-23 07:30:00
LAB | Pre-measurement calibration | pH | pH Units | Standard: pH - 7 | a. Calibration Standard | 2020-06-23 07:30:00 | 7.01
LAB | Pre-measurement calibration | pH | pH Units | Standard: pH - 7 | b. Pre-calibration Reading | 2020-06-23 07:30:00 | 7.00
LAB | Pre-measurement calibration | pH | pH Units | Standard: pH - 7 | c. Post-calibration Reading | 2020-06-23 07:30:00 | 7.01
LAB | Pre-measurement calibration | pH | pH Units | Standard: pH - 10 | a. Calibration Standard | 2020-06-23 07:30:00 | 10.04
LAB | Pre-measurement calibration | pH | pH Units | Standard: pH - 10 | b. Pre-calibration Reading | 2020-06-23 07:30:00 | 10.04
LAB | Pre-measurement calibration | pH | pH Units | Standard: pH - 10 | c. Post-calibration Reading | 2020-06-23 07:30:00 | 10.04
LAB | Pre-measurement calibration | Dissolved Oxygen | % | Standard: Dissolved Oxygen - fully saturated | a. Calibration Standard | 2020-06-23 07:30:00 | 99.62
LAB | Pre-measurement calibration | Dissolved Oxygen | % | Standard: Dissolved Oxygen - fully saturated | b. Pre-calibration Reading | 2020-06-23 07:30:00 | 100.40
LAB | Pre-measurement calibration | Dissolved Oxygen | % | Standard: Dissolved Oxygen - fully saturated | c. Post-calibration Reading | 2020-06-23 07:30:00 | 99.40
LAB | Pre-measurement calibration | Dissolved Oxygen | mg/L | Standard: Dissolved Oxygen - fully saturated | a. Calibration Standard | 2020-06-23 07:30:00 | 9.01
LAB | Pre-measurement calibration | Dissolved Oxygen | mg/L | Standard: Dissolved Oxygen - fully saturated | b. Pre-calibration Reading | 2020-06-23 07:30:00 | 9.06
LAB | Pre-measurement calibration | Dissolved Oxygen | mg/L | Standard: Dissolved Oxygen - fully saturated | c. Post-calibration Reading | 2020-06-23 07:30:00 | 9.00

## EVENT
arrival_time | cdec_code | contact_name | event_type_name | summary_name | reason_name | comments
2020-06-23 07:30:00 | LAB | PLUTON | Pre-Measurement Calibration | summary was not specified | Wagner QAQC | Verification Sonde

## ACTION
arrival_time | cdec_code | contact_name | event_type_name | sonde_name | location_name | action_name
2020-06-23 07:30:00 | LAB | PLUTON | Pre-Measurement Calibration | SM-26 | Lab | Calibrated

## INSTRUMENT
arrival_time | cdec_code | contact_name | event_type_name | instrument_name | serial_number | calibration_due_date
2020-06-23 07:30:00 | LAB | PLUTON | Pre-Measurement Calibration | Thermometer | 191957329 | 2021-02-13 00:00:00
2020-06-23 07:30:00 | LAB | PLUTON | Pre-Measurement Calibration | Turbidimeter | 2017080C0003 | 2020-07-05 00:00:00
2020-06-23 07:30:00 | LAB | PLUTON | Pre-Measurement Calibration | Barometer | 261829
2020-06-23 07:30:00 | LAB | PLUTON | Pre-Measurement Calibration
2020-06-23 07:30:00 | LAB | PLUTON | Pre-Measurement Calibration

## SOLUTION
arrival_time | cdec_code | contact_name | event_type_name | solution_name | lot_number | expiration_date
2020-06-23 07:30:00 | LAB | PLUTON | Pre-Measurement Calibration | Turbidity Standard | 21833 | 2021-01-15 00:00:00
2020-06-23 07:30:00 | LAB | PLUTON | Pre-Measurement Calibration | Conductance Standard | 20700 | 2021-11-06 00:00:00
2020-06-23 07:30:00 | LAB | PLUTON | Pre-Measurement Calibration | Conductance Check | 21860 | 2021-01-28 00:00:00
2020-06-23 07:30:00 | LAB | PLUTON | Pre-Measurement Calibration | 7 pH Standard | A0037 | 2022-02-01 00:00:00
2020-06-23 07:30:00 | LAB | PLUTON | Pre-Measurement Calibration | 4 pH Standard |  | 00:00:00
2020-06-23 07:30:00 | LAB | PLUTON | Pre-Measurement Calibration | 10 pH Standard | A0016 | 2021-01-01 00:00:00

## DOCALC
Inputs |  | Temp in Kelvin (T) | Salinity (S) | BP in atm (P) | Vapor Pressure of water in atmospheres (u) | Second Virial Coeffecient of Oxygen (0o) | DO Concentration @ zero salinity and one ATM (DO0) | Salinity Factor (Fs) | Pressure Factor (FP)
Temp in C (t) | 20.28 | 293.43 | 0 | 0.99618 | 0.023477 | 0.00071228 | 9.0421 | 1 | 0.9961
SpCond (µS/cm) | 0
Barometric Pressure (mmHg) | 757.1
|  |  | Oxygen Solubility | 9.0068 | mg/L
|  |  |  | 99.6184 | %