## ENTRY
Post-Deployment Check
Water Quality Instrument Record (SOP v1.0.0)
CA Department of Water Resources
Division of Environmental Services, Suisun Marsh Branch
[7 empty] | Sonde ID |  |  | SM-43
Discrete | Real Time |  | x | Performed by (Last Name) | [6 empty] | Moore | [5 empty] | Date Removed |  | 2020-06-16
Site Removed From / Run Used On | [5 empty] | SNC | [8 empty] | Time Removed (PST) |  |  |  | 900
Sonde Cleaned? (If applicable) |  |  |  |  | y | Attachments?
Data downloaded & Copied to Server | [5 empty] | y | Filename (include S/N)
Verification Instrument |  | Serial Number |  |  |  | Calibration due date |  |  |  |  | Calibration Standard |  |  |  |  | Lot Number |  |  | Exp. Date
[11 empty] | Turbidity Standard |  |  |  |  | 21833 |  |  | 2021-01-15
Thermometer |  | 191957329 |  |  |  | 2021-02-13 |  |  |  |  | Conductance Standard |  |  |  |  | 20700 |  |  | 2021-11-06
Turbidimeter |  | 2017080C0003 |  |  |  | 2020-07-05 |  |  |  |  | Conductance Check |  |  |  |  | 21860 |  |  | 2021-01-28
Barometer |  | 261829 | [8 empty] | 7 pH Standard |  |  |  |  | A0037 |  |  | 2022-02-01
addl. instrument | [10 empty] | 4 pH Standard
addl. instrument | [10 empty] | 10 pH Standard |  |  |  |  | A0016 |  |  | 2021-01-01
Post-deployment Check
To be completed within 24 hours of removal from deployment. Follow current standard operating procedures.
Performed by (Last Name) |  |  |  | Pluton | [7 empty] | Date |  | 2020-06-23 |  |  |  | Time (PST) | 1300
Parameter |  |  |  |  | Standard |  |  | Reading |  |  | Deviation |  | Passing Criteria |  |  |  |  | Additional Info
Specific Conductance (μS/cm) Dry |  |  |  |  | 0 |  |  | 0 |  |  | 0 |  | < 2 μS/cm
Temperature (°C) in Water |  |  |  |  | 19.94 |  |  | 20.02 |  |  | 0.08 |  | ≤ ± 0.20 °C
Chlorophyll (RFU) in DIW |  |  |  |  | 0 | [7 empty] | < 0.10 RFU
Chlorophyll (μg/L) in DIW |  |  |  |  | 0 | [7 empty] | < 0.10 μg/L
Blue Green Algae (RFU) in DIW | [12 empty] | < 0.10 RFU
Blue Green Algae (μg/L) in DIW | [12 empty] | < 0.10 μg/L
Turbidity (FNU) in DIW |  |  |  |  | 0 | [7 empty] | ≤ ± 0.5 NTU
Turbidity (FNU) in Std | [12 empty] | ≤ ± 5%
Specific Conductance (μS/cm) in Std |  |  |  |  | 17500 |  |  | 17587 |  |  | 0.4971% |  | ≤ ± 3%
pH 7 (units) | [12 empty] | ≤ ± 0.2 units | [6 empty] | mV
pH 4 (units) | [12 empty] | ≤ ± 0.2 units | [6 empty] | mV
pH 10 (units) | [12 empty] | ≤ ± 0.2 units | [6 empty] | mV
Dissolved Oxygen (% Sat) | [12 empty] | ≤ ± 5% |  |  |  |  | Baro. Pres. (mmHg):
Dissolved Oxygen (mg/L) | [12 empty] | ≤ ± 0.3 mg/L |  |  |  |  | Temp. (°C) in bucket :
Delta slope (pH 7 mV - pH 10 mV) | [8 empty] | Ideal range 160 - 180 mV, replace module if slope ≤ 155 mV
Wiper activated during check
Comments

## RESULT
cdec_code | interval_name | analyte_name | unit_name | rank_name | reading_type_name | time | value
LAB | Post-deployment check | Specific Conductance | μS/cm | Standard: Specific Conductance - Dry | a. Calibration Standard | 2020-06-23 13:00:00 | 0
LAB | Post-deployment check | Specific Conductance | μS/cm | Standard: Specific Conductance - Dry | b. Post-deployment Reading | 2020-06-23 13:00:00 | 0
LAB | Post-deployment check | Specific Conductance | μS/cm | Standard: Specific Conductance - Standard Solution | a. Calibration Standard | 2020-06-23 13:00:00 | 17500
LAB | Post-deployment check | Specific Conductance | μS/cm | Standard: Specific Conductance - Standard Solution | b. Post-deployment Reading | 2020-06-23 13:00:00 | 17587
LAB | Post-deployment check | Temperature | °C | Standard: Temperature - Water Bath | a. Calibration Standard | 2020-06-23 13:00:00 | 19.94
LAB | Post-deployment check | Temperature | °C | Standard: Temperature - Water Bath | b. Post-deployment Reading | 2020-06-23 13:00:00 | 20.02
LAB | Post-deployment check | Chlorophyll | RFU | Standard: DIW | a. Calibration Standard | 2020-06-23 13:00:00 | 0.00
LAB | Post-deployment check | Chlorophyll | RFU | Standard: DIW | b. Post-deployment Reading | 2020-06-23 13:00:00
LAB | Post-deployment check | Chlorophyll | μg/L | Standard: DIW | a. Calibration Standard | 2020-06-23 13:00:00 | 0.00
LAB | Post-deployment check | Chlorophyll | μg/L | Standard: DIW | b. Post-deployment Reading | 2020-06-23 13:00:00
LAB | Post-deployment check | Blue green algae | RFU | Standard: DIW | a. Calibration Standard | 2020-06-23 13:00:00
LAB | Post-deployment check | Blue green algae | RFU | Standard: DIW | b. Post-deployment Reading | 2020-06-23 13:00:00
LAB | Post-deployment check | Blue green algae | μg/L | Standard: DIW | a. Calibration Standard | 2020-06-23 13:00:00
LAB | Post-deployment check | Blue green algae | μg/L | Standard: DIW | b. Post-deployment Reading | 2020-06-23 13:00:00
LAB | Post-deployment check | Turbidity | NTU | Standard: DIW | a. Calibration Standard | 2020-06-23 13:00:00 | 0.00
LAB | Post-deployment check | Turbidity | NTU | Standard: DIW | b. Post-deployment Reading | 2020-06-23 13:00:00
LAB | Post-deployment check | Turbidity | NTU | Standard: Turbidity - High | a. Calibration Standard | 2020-06-23 13:00:00
LAB | Post-deployment check | Turbidity | NTU | Standard: Turbidity - High | b. Post-deployment Reading | 2020-06-23 13:00:00
LAB | Post-deployment check | pH | pH Units | Standard: pH - 4 | a. Calibration Standard | 2020-06-23 13:00:00
LAB | Post-deployment check | pH | pH Units | Standard: pH - 4 | b. Post-deployment Reading | 2020-06-23 13:00:00
LAB | Post-deployment check | pH | pH Units | Standard: pH - 7 | a. Calibration Standard | 2020-06-23 13:00:00
LAB | Post-deployment check | pH | pH Units | Standard: pH - 7 | b. Post-deployment Reading | 2020-06-23 13:00:00
LAB | Post-deployment check | pH | pH Units | Standard: pH - 10 | a. Calibration Standard | 2020-06-23 13:00:00
LAB | Post-deployment check | pH | pH Units | Standard: pH - 10 | b. Post-deployment Reading | 2020-06-23 13:00:00
LAB | Post-deployment check | Dissolved Oxygen | % | Standard: Dissolved Oxygen - fully saturated | a. Calibration Standard | 2020-06-23 13:00:00
LAB | Post-deployment check | Dissolved Oxygen | % | Standard: Dissolved Oxygen - fully saturated | b. Post-deployment Reading | 2020-06-23 13:00:00
LAB | Post-deployment check | Dissolved Oxygen | mg/L | Standard: Dissolved Oxygen - fully saturated | a. Calibration Standard | 2020-06-23 13:00:00
LAB | Post-deployment check | Dissolved Oxygen | mg/L | Standard: Dissolved Oxygen - fully saturated | b. Post-deployment Reading | 2020-06-23 13:00:00

## EVENT
arrival_time | cdec_code | contact_name | event_type_name | summary_name | reason_name | comments
2020-06-23 13:00:00 | LAB | PLUTON | Post-Deployment Check | summary was not specified | Wagner QAQC | No comments.

## ACTION
arrival_time | cdec_code | event_type_name | contact_name | sonde_name | location_name | action_name
2020-06-23 13:00:00 | LAB | Post-Deployment Check | PLUTON | SM-43 | Lab | Checked

## INSTRUMENT
arrival_time | cdec_code | contact_name | event_type_name | instrument_name | serial_number | calibration_due_date
2020-06-23 13:00:00 | LAB | PLUTON | Post-Deployment Check | Thermometer | 191957329 | 2021-02-13 00:00:00
2020-06-23 13:00:00 | LAB | PLUTON | Post-Deployment Check | Turbidimeter | 2017080C0003 | 2020-07-05 00:00:00
2020-06-23 13:00:00 | LAB | PLUTON | Post-Deployment Check | Barometer | 261829
2020-06-23 13:00:00 | LAB | PLUTON | Post-Deployment Check
2020-06-23 13:00:00 | LAB | PLUTON | Post-Deployment Check

## SOLUTION
arrival_time | cdec_code | contact_name | event_type_name | solution_name | lot_number | expiration_date
2020-06-23 13:00:00 | LAB | PLUTON | Post-Deployment Check | Turbidity Standard | 21833 | 2021-01-15 00:00:00
2020-06-23 13:00:00 | LAB | PLUTON | Post-Deployment Check | Conductance Standard | 20700 | 2021-11-06 00:00:00
2020-06-23 13:00:00 | LAB | PLUTON | Post-Deployment Check | Conductance Check | 21860 | 2021-01-28 00:00:00
2020-06-23 13:00:00 | LAB | PLUTON | Post-Deployment Check | 7 pH Standard | A0037 | 2022-02-01 00:00:00
2020-06-23 13:00:00 | LAB | PLUTON | Post-Deployment Check | 4 pH Standard
2020-06-23 13:00:00 | LAB | PLUTON | Post-Deployment Check | 10 pH Standard | A0016 | 2021-01-01 00:00:00

## DOCALC
Inputs |  | Temp in Kelvin (T) | Salinity (S) | BP in atm (P) | Vapor Pressure of water in atmospheres (u) | Second Virial Coeffecient of Oxygen (0o) | DO Concentration @ zero salinity and one ATM (DO0) | Salinity Factor (Fs) | Pressure Factor (FP)
Temp in C (t) | 0 | 273.15 | 0 | 0 | 0.0060266 | 0.000975 | 14.6208 | 1 | -0.0060691
SpCond (µS/cm) | 0
Barometric Pressure (mmHg) | 0
|  |  | Oxygen Solubility | -0.088735 | mg/L
|  |  |  | 0 | %